## Annex A. Financial Bid
|  | ITB reference number: ITB-POL-00012 - Medical and Therapeutic Equipment |  |  |  |  | Annex A.1 Financial Bid
DRC to complete |  |  |  | Bidder to complete
# | Item Required | Specification | Quantity required | Item offered (name, make, and model, with full specification, warranty duration) | Quantity offered | Unit Price (with VAT) | Total Price (with VAT)
Lot 1
1 | Hot-tub medical for hydro-massage | Therapeutic bathtub for zone hydromassage, economical use of water - min. 160 L, max. 320 L; 26 nozzles for hydro massage divided into 3 separate zones (8, 8, 10 nozzles), 126 aeration nozzles to air bubble massage, ozone generator, independent air pump, automatic filling, pump "dry-run" protect., disinfection and descaling system, system of removing water from the water ducts, colour touch panel, 2 temp. sensors (measurement while filling the bat.h and during treatment), maintains water temp. automatically (heater), warm and cold water shower, headrest, steps. Minimum of 2 years seller warranty period. Installation of the equipment to be inclusive in the price. | 1
2 | Whirl massage device for upper limbs | Whirl massage device for upper limbs, economical use of water – min. 30 L, max. 45 L, 44 aeration nozzles for massage, pump "dry-run" protection, electronically controlled descaling system, semi-automatic draining system, electronic control panel, temperature sensor, water level sensor, warm shower, chair. Minimum of 2 years seller warranty period. Installation of the equipment to be inclusive in the price. | 1
3 | Water pump for massage Aquavibron | Water supply station, with a closed water circulation system, supplying AQUAVIBRON vibration massage apparatus, 7-10 L capacity, eliminates water consumption - once filled up, does not require connection to the water supply system. Minimum of 2 years seller warranty period. Installation of the equipment to be inclusive in the price. | 1
Lot 2
1 | Polarized lamp for procedures | Wavelength 350 - 3400 nm Degree of polarization >95% (590 - 1550 nm) Specific power density av. 40 mW/cm2 Light energy per minute av. 2.4 J/cm2 Light Intensity min. 10.000 lux CE labeling CE0197 Weight without stand 3.4 kg Weight with stand 7.8 kg Energy consumption in sleep mode 0.5 W Voltage 100-240 V~ 50/60 Hz (bioptron). Minimum of 1 year seller warranty period. | 2
2 | An infrared irradiation lamp | An infrared irradiation lamp on adjustable stand, 3 infrared radiators of max. 3 x 375 W power, electronic control panel, radiator power control, treatment duration control, lamp arm sliding up / down, low-profile wheeled base for easy moving and approaching under the TERAPEUTA massage tables (system of 4 castors with breaks), protective glasses for the patient and therapist included. Minimum of 1 year seller warranty period. | 1
Lot 3
1 | Crio Therapy Machine | · Power supply: 230 V/50 Hz · Cooling medium: Carbon dioxide (CO2) in siphon cylinder · Operationg pressure: normal: 5÷6 MPa, maximum 8 MPa · Operation (cooling) time: · 10 litre (7.5 kg) cylinders – nozzle no. 3: 45 min · 10 litre (7.5 kg) cylinders – nozzle no. 1: 120 min · Nozzle temperature range: -65º ÷ -75º · 2 nozzle connection ports with (electric and pneumatic each) · LCD screen · Control panel · Sound notrification system · Dimensions: 360 x 410 x 135 (mm) · Weight: 6 kg Minimum of 2 years seller warranty period. | 1
Lot 4
1 | M-S3: 3-section massage table | M-S3: 3-section massage table, powder coated steel high/low height adjustment foot bar accessible from each side of table, easily adjustable head- and foot-rest with gas spring, 2 plugs: round and crescent F0: the base with 2 wheels and 2 legs. Width: 69 cm Minimum of 1 year seller warranty period. | 2
2 | Therapeutic examination table | 2-section examination table with long head section supporting the sitting position; powder coated steel high/low height adjustment foot bar accessible from each side of table, easily adjustable head-rest with gas spring. Width: 69 cm. F0: the base with 2 wheels and 2 legs. Minimum of 1 year seller warranty period. | 2
3 | Procedure beds | Metal medical couch - face hole, no plug. with paper holder. Minimum of 1 year seller warranty period. | 6
Sub-total
Any other costs (please specify)
Total Price
DRC to complete |  |  |  | Bidder to complete
Delivery time required (days after PO signature): |  | 20 Days |  | Delivery time offered (days after PO signature):
Delivery Terms required (Add Incoterm if necessary): |  | DDP |  | Delivery Terms offered (must include incoterm):
Delivery Destination required: |  | CRR Dobry Brat, Osiedle Dobry Brat 12, 83-221 Osiek, Poland |  | Delivery Destination offered:
Minimum bid validity period required: |  | 30 Days |  | Bid validity period offfered:
Currency of Tender: |  | PLN |  | Currency of Bid:
Additional comments to the Bidders: |  |  |  | Company Name:
Signed by a duly authorized company representative:
Title:
Date:
Print Name:
Stamp of company